## Slide 1: Vulcan Expert+ Calibration datasheet
December 2020

## Slide 2: Aluminium and Magnesium Alloys
Aluminium alloys |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Element | Be | B | Mg | Al | Si | Ti | Cr | Mn | Fe | Ni | Cu | Zn | Zr | Ag | Pb | Bi
Calibration range | 0 - 0.24 | 0 - 10 | 0 - 10.2 | 79.6-100 | 0 - 18.7 | 0 - 0.33 | 0 - 0.16 | 0 - 1.3 | 0 - 1.1 | 0 - 2.4 | 0 - 8.0 | 0 - 8.0 | 0 - 0.3 | 0 - 0.31 | 0 - 1.1 | 0.2 - 1.3
Error, wt. % | 0.005 | 1.06 | 0.26 | 1.04 | 0.79 | 0.04 | 0.05 | 0.07 | 0.11 | 0.07 | 0.23 | 0.14 | 0.04 | 0.03 | 0.13 | 0.19
Magnesium alloys |  |  |  |  | 
Element | Mg | Al | Mn | Cu | Zn
Calibration range | 82.4 - 100 | 0 - 12.4 | 0 - 1.8 | 0 - 2.9 | 0 - 6.8
Error, wt. % | 0.78 | 0.46 | 0.13 | 0.07 | 0.41
Error %: Typical difference between measured value and certified reference value when large set of test samples are measured with multiple production instruments. Notice that performance of individual instrument may differ slightly.
Calibration range: Concentration range of the standards used in calibration.

## Slide 3: Stainless Steels, Low Alloy Steels and Tool Steels
Stainless steels |  |  |  |  |  |  |  |  |  |  |  |  | 
Element | Al | Si | Ti | V | Cr | Mn | Fe | Co | Ni | Cu | Nb | Mo | W
Calibration range | 0 - 1.2 | 0 - 4.0 | 0 - 2.2 | 0 - 3.0 | 0 - 27.4 | 0 - 9.7 | 36 - 88.7 | 0 - 18.7 | 0 - 37.4 | 0 -3.4 | 0 - 1.5 | 0 - 6.5 | 0 - 3.0
Error, wt. % | 0.04 | 0.15 | 0.04 | 0.06 | 0.49 | 0.12 | 0.93 | 0.24 | 0.53 | 0.06 | 0.12 | 0.22 | 0.17
Low alloy steels |  |  |  |  |  |  |  |  |  |  |  | 
Element | Al | Si | Ti | V | Cr | Mn | Fe | Ni | Cu | Zr | Nb | Mo
Calibration range | 0 - 1.3 | 0 - 4.0 | 0 - 0.3 | 0 - 1.8 | 0 - 8.8 | 0 - 2.2 | 84.4 - 100 | 0 - 5.3 | 0 - 0.6 | 0 - 0.41 | 0 - 0.3 | 0 - 1.4
Error, wt. % | 0.01 | 0.18 | 0.02 | 0.04 | 0.17 | 0.12 | 1.2 | 0.10 | 0.03 | 0.04 | 0.02 | 0.09
Tool steels |  |  |  |  |  |  |  |  |  |  |  | 
Element | Al | Si | Ti | V | Cr | Mn | Fe | Co | Ni | Cu | Mo | W
Calibration range | 0 - 0.9 | 0 - 2.6 | 0 – 0.47 | 0 - 4.9 | 0 - 12.0 | 0 - 2.3 | 69.6 - 96 | 0 - 8.0 | 0 - 2.0 | 0 - 0.7 | 0 - 9.4 | 0 - 18.1
Error, wt. % | 0.04 | 0.20 | 0.06 | 0.20 | 0.40 | 0.06 | 2.56 | 0.18 | 0.14 | 0.03 | 0.31 | 1.11
Error %: Typical difference between measured value and certified reference value when large set of test samples are measured with multiple production instruments. Notice that performance of individual instrument may differ slightly.
Calibration range: Concentration range of the standards used in calibration.

## Slide 4: Nickel, Copper and Titanium Alloys
Nickel alloys |  |  |  |  |  |  |  |  |  |  |  |  | 
Element | Al | Si | Ti | Cr | Mn | Fe | Co | Ni | Cu | Nb | Mo | Hf | W
Calibration range | 0 - 6.0 | 0 - 3.9 | 0 - 5.2 | 0 - 29.5 | 0 - 1.6 | 0 - 42.7 | 0 - 19.0 | 33.3 - 100 | 0 - 33.0 | 0 - 5.1 | 0 - 26.5 | 0 - 1.5 | 0 - 14.1
Error, wt. % | 0.29 | 0.38 | 0.20 | 0.99 | 0.08 | 1.4 | 0.64 | 2.5 | 0.49 | 0.27 | 1.1 | 0.12 | 0.72
Copper alloys |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Element | Be | Mg | Al | Si | Cr | Mn | Fe | Co | Ni | Cu | Zn | Ag | Sn | Pb | Bi
Calibration range | 0 - 1.9 | 0 - 0.19 | 0 - 13.5 | 0 - 3.6 | 0 - 0.8 | 0 - 3.0 | 0 - 5.4 | 0 - 2.3 | 0 - 30.2 | 57.9 - 100 | 0 - 39.3 | 0 - 0.7 | 0 - 12.6 | 0 - 11.7 | 0 - 6.5
Error, wt. % | 0.02 | 0.01 | 0.22 | 0.08 | 0.03 | 0.06 | 0.04 | 0.09 | 0.36 | 1.4 | 0.88 | 0.02 | 0.35 | 0.66 | 0.18
Titanium alloys |  |  |  |  |  |  |  |  |  | 
Element | Al | Ti | V | Cr | Mn | Fe | Zr | Nb | Mo | Sn
Calibration range | 0 - 6.3 | 74.8 - 100 | 0 - 15.3 | 0 - 6.3 | 0 - 2.0 | 0 - 1.2 | 0 - 4.1 | 0 - 6.9 | 0 - 6.2 | 0 - 3.1
Error, wt. % | 0.27 | 2.9 | 0.18 | 0.28 | 0.07 | 0.13 | 0.13 | 0.47 | 0.26 | 0.46
Error %: Typical difference between measured value and certified reference value when large set of test samples are measured with multiple production instruments. Notice that performance of individual instrument may differ slightly.
Calibration range: Concentration range of the standards used in calibration.

## Slide 5: Zinc and Cobalt Alloys
Zinc alloys |  |  |  |  |  | 
Element | Al | Fe | Cu | Zn | Sn | Pb
Calibration range | 0 - 4.5 | 0 - 0.5 | 0 - 1.5 | 94.7 - 100 | 0 - 0.15 | 0 - 1.0
Error, wt. % | 0.12 | 0.08 | 0.01 | 1.2 | 0.03 | 0.04
Cobalt alloys |  |  |  |  |  |  |  |  |  | 
Element | Si | Ti | Cr | Mn | Fe | Co | Ni | Nb | Mo | W
Calibration range | 0 – 1.2 | 0 - 3.0 | 19 - 31 | 0 - 1.9 | 0 - 15.2 | 24 - 65 | 0.1 - 25.7 | 0 - 2.4 | 0 - 7.9 | 0 - 15.1
Error, wt. % | 0.25 | 0.05 | 1.4 | 0.08 | 0.17 | 1.9 | 0.4 | 0.18 | 0.3 | 0.99
Error %: Typical difference between measured value and certified reference value when large set of test samples are measured with multiple production instruments. Notice that performance of individual instrument may differ slightly.
Calibration range: Concentration range of the standards used in calibration.

## Slide 6: Lead, Tin and Lead+Tin Alloys
Lead alloys |  |  | 
Element | Sn | Sb | Pb
Calibration range | 0 - 10.1 | 0 - 9.6 | 89.9 - 100
Error, wt. % | 0.18 | 0.24 | 0.32
Tin alloys |  |  |  |  |  | 
Element | Cu | Ag | Sn | Sb | Pb | Bi
Calibration range | 0 - 8.0 | 0 - 4.0 | 83.1 - 100 | 0 - 8.3 | 0 - 0.3 | 0 - 1.1
Error, wt. % | 0.62 | 0.09 | 0.80 | 0.18 | 0.03 | 0.03
Lead + Tin alloys |  |  | 
Element | Ag | Sn | Pb
Calibration range | 0 - 2.0 | 30.9 - 62.4 | 35.7 - 68.4
Error, wt. % | 0.05 | 1.46 | 1.27
Error %: Typical difference between measured value and certified reference value when large set of test samples are measured with multiple production instruments. Notice that performance of individual instrument may differ slightly.
Calibration range: Concentration range of the standards used in calibration.